## Slide 1
Introduction to Proteomics
Roz Banks & Alexandre Zougman
(r.banks@leeds.ac.uk) (a.zougman@leeds.ac.uk)
Clinical and Biomedical Proteomics Group

## Slide 2
Clinical and Biomedical Proteomics Group
Part of the Biomarkers and Therapy Section in the Leeds Institute of Cancer and Pathology within Faculty of Medicine and Health
Relatively small multidisciplinary group (n=6 currently)
Focussed on using proteomics technologies for biomarker discovery/translation (+ mechanistic insights and novel therapeutic targets…)
Provide a local research facility for proteomics studies (www.proteomics.leeds.ac.uk)
Complements the much larger mass spectrometry facility provided in FBS (www.fbs.leeds.ac.uk/facilities/ms) – focus on protein structure and conformation

## Slide 3
What is Proteomics?
“Proteome" – first used in 1994 by Marc Wilkins through analogy with the genome. The entire complement of proteins, including the modifications made to a particular set of proteins, produced by an organism or a cellular system.
“Proteomics“ – the large-scale study of a specific proteome. Can include many aspects such as protein abundances, modifications, sub-cellular localisation and interacting partners and networks, in order to understand cellular processes.
Proteins represent the “functional molecules” of the cells or organism

## Slide 4
Critical Considerations/Challenges
One gene can encode more than one protein.
Human genome ~21,000 protein-encoding genes.
Added complexity - with co- and post-translational modifications, human proteins estimated 250,000 to one million.
Implications for Sample Preparation – Pre-fractionation – “Snapshots”
Proteomes are highly dynamic compared with the relatively static genome. Varies over time and in response to stimuli e.g. stress, disease etc – amounts or PTM forms of proteins.
Implications for Sample Preparation – Stability
Proteins exist in a wide range of concentrations. For example, the concentration of the protein albumin in blood is more than a billion times greater than that of interleukin-6 and 20 proteins make up >98% of the serum proteome.
Implications for Sample Preparation – Pre-fractionation (no protein equivalent of PCR)

## Slide 5
Types of Proteomics
Clinical Proteomics - use of proteomics to address clinical questions such as biomarker discovery or pathways involved in disease
Glycoproteomics
Phosphoproteomics
Chemical Proteomics - using chemical tools to study the modification or interaction of small molecules with proteins
Variously named by scientific context or focus
Examples:
Bead
e.g. drug target
identification
Direct affinity resin enrichment of proteins that bind to a drug
Drug is linked to beads
Protein lysate is incubated with beads
Bound proteins identified by mass spectrometry

## Slide 6
Proteomic Technologies 1
200
-VHL
+VHL
200
14
kDa
pI 4 7
Two-dimensional polyacrylamide
gel electrophoresis
Not used very often now
~2,000 proteins visualised – very good
for seeing multiple forms
But:
Very limited for hydrophobic proteins such as membrane proteins
Multiple proteins in one spot
No immediate identification – use MS
Limited dynamic range

## Slide 7
Array-based techniques
e.g. antibodies, phage or reverse
phase arrays, multiplex ELISA
Targeted so determined by specific binders
Typically 50 to 2-3,000 proteins
Mass Spectrometry
Now regarded by many as the workhorse of proteomics
Main focus of the rest of this talk
Proteomic Technologies 2
Example

## Slide 8
Label-free profiling
LC-MS/MS (DDA)
Thermo Velos
Fractionation e.g.
immunodepletion
F N N V S
Bioinformatic &
statistical analysis
Relative quantification of
~450 serum proteins
~3-4,000 tissue proteins
~1,200 urinary proteins
“Typical” shotgun approach in Leeds
Tissue
Cell lines
Blood
Urine
Study Design & Sample Selection

## Slide 9
Critical Aspects of Proteomic Studies
Well defined clinical/scientific question and study population/system
Choice of technology and approach – scientific input
Appropriate study design
- avoidance of bias/confounding factors
- sample size and replicates (knowledge of technical and
biological variance), statistical input
Sample Preparation
- e.g. extraction buffer composition and additives, chemical quality
- pre-fractionation
- containers e.g. type of blood tubes, plastics etc
Data Analysis
Validation approach
- e.g. western blotting, IHC, ELISA etc
N
T
N
T
N
T
N
T
N
T

## Slide 10: Proteomics and Mass Spectrometry
making sense out of fragments

## Slide 11
Proteomics is an art of identification, characterization and quantitation of proteins or complete proteomes (protein pools)
“Tout progrès scientific est un progrès de method.”
“All scientific progress is progress of a method.
René Descartes
Mass spectrometry (MS) is a vital tool of proteomics
During a typical proteomics experiment, the proteins are digested by proteases, the resultant peptide mixtures are separated and introduced into a mass spectrometer by a liquid chromatography (LC) system, and the distinct peptide signatures - the exact peptide mass and partial internal amino acid composition - are deduced from MS and tandem mass spectrometry (MS/MS) data
The examples illustrating the use of mass spectrometry in protein research will be presented

## Slide 12
Mass spectrometers use the difference in mass-to-charge ratio (m/z) of ionized molecules for separation and detection
M/Z

## Slide 13: It is always Mass over Charge
H +
H +
H +
H +
H +
H +

## Slide 14
Sample Prep
Capillary HPLC
Mass Spectrometer
Data processing and presentation

## Slide 15
UHPLC
High Accuracy Mass Spectrometer
Dionex UHPLC + LTQ Orbitrap Velos Mass Spectrometer

## Slide 16
Protein
Digestion
Bottom-Up
Peptide sequencing
Peptide /
Protein identification
Protein sequencing
Protein identification
Peptides
Top-Down
Proteomic Approaches

## Slide 17
Protein digestion is a prerequisite of the MS analysis

## Slide 18
In a typical proteomics experiment a protein mixture is digested by an enzyme thus creating a mixture of peptides. The most common protease is trypsin which cuts protein C-terminally at arginine (R) or lysine (K) residues and gives rise to peptides carrying at least one basic residue. At pH 3, Asp- and Glu- are uncharged and any proteolytic peptide has the net charge of at least 2+ where one charge is located C-terminally at the basic residue and another charge is localized at the N-terminus.
The opposite location of the charges facilitates peptide fragmentation at the amide bond in the most commonly used low energy collision induced dissociation (CID) process creating y and b ion series according to Roepstorff-Fohlmann-Biemann nomenclature.

## Slide 19
Nano-LC interface

## Slide 20
Columns utilizing arch types found in stone-bridges. (a) Tapered column with the ‘‘keystone’’ arch, (b) tapered column with the ‘‘double-stone’’ arch; (left picture) a stone-bridge with the keystone arch, (right picture) a stone-bridge with the double-stone arch (Ishihama et al. J Chromatogr A 979, 233 (Dec 6, 2002).
75 µm (ID) ~ 6 µm tip
Fused silica spraying capillary packed
with 3µm C18 silica particles.
X400 magnification.
Spraying capillary column

## Slide 21
Gradient elution:200 nl/min
Column (75 µm)/spray tip (6 µm)
Reverse-phase C18 beads, 3 µm
Platinum-wire
2.0 kV
Sample Loading:500 nl/min
No precolumn or split
MS
15 cm
Nanoflow LC/MS interface set-up
Coupling LC to MS for complex mixture analysis

## Slide 22
Taylor cone
Capillary with
analyte
Plume
Electrospray Ionization (ESI) process introduces ionized molecules into a mass spectrometer

## Slide 23
The Orbitrap analyzer is the example of how the old concept if resurrected at the right time by an exceptional inventor could lead to creation of a truly revolutionary device. Orbital trapping was originally introduced by Kingdon in 1923.
The Orbitrap Mass Spectrometer
However, the estimation of the m/z ratios was based on the frequencies of ion rotations. The approach lead to poor mass resolution because ion velocity and initial radius significantly influence the rotational frequency.
The axial frequency is detected by processing the image current with Fourier Transform (FT) algorithms.
Makarov’s idea was to derive the m/z values from the frequency of harmonic ion oscillations along the field’s axis - this frequency is completely independent of energy and spatial spread of ions (patented in 1999).
Dr. Alexander A. Makarov joined a small high-tech start-up company called HD Technologies (Manchester, UK) in the mid-1990s.

## Slide 24
Axial oscillation
The Orbitrap

## Slide 25: Linear Ion Trap – Orbitrap Hybrid Mass Spectrometer (LTQ-Orbitrap)
1. Ions are stored in the Linear Trap
2. …. are axially ejected
3. …. and trapped in the C-trap
4. …. they are squeezed into a small cloud and injected into the Orbitrap
5. …. where they are electrostatically trapped, while rotating around the central electrode and performing axial oscillation
The oscillating ions induce and image current into the two outer halves of the orbitrap, which can be detected using a differential amplifier
Ions of only one mass generate a sine wave signal

## Slide 26: Frequencies and Masses
The axial oscillation frequency follows the formula
Where w = oscillation frequency
k = instrumental constant
m/z = …. well, we have seen this before
Many ions in the Orbitrap generate a complex signal whose frequencies are determined using a Fourier Transformation

## Slide 27
Most chemical elements exist as diverse masses in nature, the isotopes. Amino acids, the building units of proteins, contain carbon (C), hydrogen (H), nitrogen (N), oxygen (O), sulfur (S) and phosphorus (P).
About 1.11 % of the naturally occurring carbon C13 atoms are 1.0034 Da heavier than the C12 atoms. It means that from a random set of 100 carbon atoms about 99% will have a mass of 12.0000 Da and 1% will have a mass of 13.0034 Da.

## Slide 28
The bigger the polypeptide mass, the higher is the probability of C13 incorporation into this molecule and the more prominent is the isotopic effect.
Monoisotopic peak

## Slide 29
M/Z

## Slide 30
Full MS scan is performed to choose the ions for MS/MS

## Slide 31
0.5
0.5
2 +
The mass is known
Precursor ion is chosen for fragmentation

## Slide 32
S
V
N
F
N
The fragmentation signature of the parent ion
MS/MS fragmentation of the precursor 555.793 2+ ion

## Slide 33
Mass of the peptide AND its fragmentation pattern = identification
VFSVNNFQR peptide of GPR 158

## Slide 34
Mass of the peptide AND its fragmentation pattern = identification
ONLY if the protein sequence is known and present in the queried database

## Slide 35
GPR158 coverage map

## Slide 36
The mass accuracy of a mass spectrometer is defined as the difference between the measured mass and its calculated value. Traditionally relative mass accuracy is reported and defined in parts per million (ppm) – the mass accuracy divided by a calculated value of the measured mass.
High mass accuracy is directly related to the ability of a mass spectrometer to separate the adjacent peaks, to its resolution. The resolution could be calculated by dividing m/z of a peak by its width at a certain height (m/z divided by Δm/z). A common definition of resolution is Full Width of the peak at Half of its Maximum height, abbreviated as 'FWHM', the Δm/z value in this case is taken at 50% of the peak height .
Resolution 4000
Resolution 50000

## Slide 37
Tryptic HSP-70 peptide: ELEEIVQPIISK, mass 1396.7813 Da
Instrument | LTQ | QSTAR | LTQ-FT | LTQ-Orbitrap
Mass Accuracy | 500 ppm | 20 ppm | 2 ppm | 1 ppm
# of tryptic peptides for m/z 1396.7813 | 344 | 52 | 11 | 9
High-mass accuracy – what is it good for?

## Slide 38
Most proteins are modified
The modifications could be accidental (e.g. Met oxidation), artificial (e.g. Cys derivatisation ) or natural (post-translational, e.g. phosphorylation)
When a modification is introduced, we observe a change in the mass of a peptide/protein carrying the modified amino acid – delta mass
serine
phosphoserine
+ 79.966 Da
Delta mass

## Slide 39
N
CH
H
C
O
(CH2)3
H2C
N
H3C
CH3
Nε-dimethyl
lysine
Mass difference = 0.03638 Da
Well-known PTM: Dimethylation
Nominal mass:
28 Da
High Mass Accuracy Enables Discovery of Novel Modifications
N
CH
H
C
O
(CH2)3
H2C
NH
O C
H
Nε-formyl
lysine
Wiśniewski JR, Zougman A, Mann M. N-epsilon-formylation of lysine is a widespread post-translational modification
of nuclear proteins occurring at residues involved in regulation of chromatin function. Nucleic Acids Res. 2008
Feb;36(2):570-7.
n=24 n=27
Formylation

## Slide 40
Gel based 2D PAGE (DIGE)
Label-free
Labeling with stable isotopes
13C, 15N, 2H
In vivo metabolic labeling (SILAC)
Affinity tags
ICAT, HysTag
Isobaric tags
iTRAQ
Chemical labeling
Mass spectrometer
based
MS1
MS2
Quantitation methods in proteomics

## Slide 41
Advantage:
High dynamic range
Coverage of entire proteomes (soluble and membrane proteins)
Weakness:
Not applicable to the real-life samples (e.g. clinical material)
Expensive
Upregulated
protein - Peptide ratio >1
Background protein - Peptide ratio 1:1
Light Heavy
SILAC relies on metabolic incorporation of a given 'light' or 'heavy' form of the amino acid into the proteins.
Stable isotope labeling with amino acids in cell culture
(SILAC)

## Slide 42
State A
State B
Proteins Identification and Quantitation by LC-MS
Protein extraction
Protein Identification and Quantitation by LC-MS
Run A
Run B
Upregulated
protein - Peptide ratio >1
Background protein - Peptide ratio 1:1
Label-free proteome quantitation

## Slide 43
Suspension trapping (STrap) sample preparation method for bottom-up proteomics analysis
Proteomics. 2014 May;14(9):1006-0. Zougman A, Selby PJ, Banks RE
Sodium Dodecyl Sulfate
Removal of detergents is the key...
SDS
Detergents are the reagents of choice for solubilization
of proteins
Due to their ready ionization ability and high abundance detergents are incompatible with ESI MS analysis of peptides
Removal of detergents is a prerequisite for successful LC-MS profiling of peptide mixtures

## Slide 44
Basic STrap method
Detergent depletion and protein digest is done in a pipette tip
Up to 50 µg of protein could be efficiently processed in the STrap tip
Depletion of SDS and introduction of a protease
takes ca. 5 min
The digest is performed in 10-15 µl buffer solution. Tryptic digest is complete in 60 min at 47°C
STrap is compatible with even elevated SDS concentrations (up to ca. 15%)
It is a simple method which does not require
sophisticated instrumentation and is aimed at high-throughput comprehensive sample processing
The basic STrap method links SDS-based protein extraction with rapid detergent removal, reactor-type protein digestion and peptide clean-up
SDS-solubilized proteins are acidified and introduced into the STrap-tip with the quartz depth filter and C18 compartments, filled with the methanolic solution. The fine protein suspension is trapped in the depth filter. SDS and other contaminants are removed in the flow-through, and a protease is introduced. Following the digestion, the peptides are cleaned-up in the tip’s C18 part.
The clean-up of the peptides takes ca. 5 min

## Slide 45
200
116
66
31
22
14
7
MW, kDa
HeLa lysate
Flow-through
Digest
Leftover
200
116
66
31
22
14
7
MW, kDa
HeLa lysate
30 min
6 min
15 min
3 min
2295
2530
2802
3500
Proteins identified
cytoplasm
nucleus
membrane
Annotated (%)
Our dataset
Dataset of Nagaraj et al., 2011
STrap and cell lysates

## Slide 46
Some Examples from Proteomics Studies Using Mass Spectrometry

## Slide 47
4 Normal/Tumour Frozen Tissue Pairs
Sample
prep
LC-MS/MS VELOS runs
MaxQuant label-free analysis
N
T

## Slide 48
0
8.25
2.46
1.07
0.53
886 proteins
Total Intensity Ratio Tumor/Normal
100
5355 proteins
SLC5A1
SLC7A8
CD11C
CD11B
CD41
CD276
Claudin-1
EPHB4
RAI3
Intracellular
N T
N T
N T
N T
1
2
3
4
RAI3
Lamin A/C
Identification and confirmation of RAI3 overexpression in malignant versus normal tissue lysates.
Cumulative label free intensity ratios for the identified normal and malignant colon proteins. The names of plasma membrane proteins with the intensity ratios exceeding 3xIQR confirmed in at least three pairs are displayed.
Western Blot analysis of RAI3 expression in matched normal (N)/tumor (T) colon tissue pairs. Mirroring the LC-MS/MS quantitation data, the blot confirms RAI3 overexpression in all four cancer samples. Lamin A/C is used as a reference.

## Slide 49
B
A
C
D
Examples of RAI3 luminal and intracytoplasmic vesicular/granular staining in colon cancer and in normal colon epithelium. (A) Luminal staining in colon cancer (arrowed); (B) Intra-cytoplasmic vesicles/granules in colon cancer (arrowed); (C) Negative staining, normal epithelium with weak stromal signal. Note the neuroendocrine cells (arrowed); (D) High power view of neuroendocrine cells staining in normal colonic crypts.

## Slide 50
A
B
C
D
Example staining intensities of RAI3 in colon cancer. (A) Negative; (B) Weak epithelial signal intensity with strong stromal signal; (C) Moderate intensity; (D) Strong intensity.

## Slide 51
Kaplan–Meier plot of recurrence-free survival for stage I-III colon cancer patients based on dichotomized cytoplasmic RAI3 expression in tumor epithelium. RAI3 epithelial cytoplasmic expression is dichotomized as negative, weak and moderate versus strong.

## Slide 52
Immunodepleted of 14 most abundant proteins
LC-MS/MS (VELOS)
(label-free
quantitation)
5 DGF patients & 5 non-DGF
(20 serum samples: Pre- and d2 post-transplant)
Matched for clinical variables
Aminoacylase-1 (ACY-1)
Renal Transplantation
Biomarkers for Earlier Detection of Delayed Graft Function and Outcome

## Slide 53
b.
ELISA Developed
Day 1 Serum ACY-1
measured
b.
Cohort 1 (n=47)
Cohort 2 (n=194)

## Slide 54
Prognostic value of serum ACY-1 d1 or 3
post-transplant in DGF patients
Welberry Smith et al. Kidney Int. 84:1214-25 (2013)
(n=52)